## Guidance notes
Guidance notes
Update September 2021:
This update to the ODEP form is intended to ensure that the panel is fully aware of the nature and scope of variations in design and materials within a brand (or brand variant) submitted for a rating.
Where clinical data is available for different variants within a brand, manufacturers are encouraged to make separate submissions with the associated data for the specified set of product codes. However, we appreciate that this may not be possible with some sources of clinical data. In those cases, full disclosure of which variants are included in the clinical data is essential for a rating to be applied.
It has always been ODEP’s policy to encourage manufacturers to analyse the clinical data available to them on their products. Now that increasingly detailed data is available through registries you will be strongly encouraged to analyse this for the purposes of a rating.
Ratings are in line with the ODEP ratings system shown on the ratings system tab
1 | Complete EITHER "Prosthesis Fields - Stem" OR "Prosthesis Fields - Cup" as appropriate
2 | Please see the website - ODEP Methodology for Hips
3 | ODEP will only accept registry data which has been published in the latest data available from that registry. Where data from publications and public presentations is used, it will only be accepted within 2 years of publication. For in-house validated data, ODEP will only accept contemporaneous data. ODEP may request manufacturers to validate any in-house data that they use.
4 | Submissions for products where the design has changed should only include data for the current variant.
5 | Ratings will only be awarded where ALL requested data fields are completed in full on ALL sheets.
6 | Each clinical reference requires a separate clinical data sheet to be submitted
7 | ODEP does not benchmark hemiarthroplasties
8 | At present ODEP does not benchmark implants which are specifically used in a revision situation.
9 | The yellow boxes on the prosthesis fields, when they are clicked, will give you drop-down options for your answers.
Data submitted must be representative of all data available. Any variation should be highlighted on the Prosthesis Fields tab

## Prosthesis Fields - Stem
Hip Prosthesis Data - Stem
Please read the guidance notes before filling out this sheet
Prosthesis details
Manufacturer
UK Distributor (if different)
Stem brand (please name all variants or versions within the Brand)
Product Codes and Sizes |  | Please provide on a separate sheet detailing individual product numbers and descriptions in Excel format
Images of products attached, including all variants |  | Photo taken against white background, best size 300x300px, resolution 200dpi. This could be used on the ODEP website.
The date (year) of the first ODEP award for this product
Date of this submission
Please confirm for the product codes listed
Data presented relates to use in primary hip replacement only
Material composition is identical
Surface finish and coating is identical
Taper / trunnion design is identical
Proximal Geometry
If present, material of modular neck is identical
If you answer 'no' to any of the above please explain and ensure it is clear what differences exist and the product codes they relate to
Technical design features for the product codes listed
Fixation
If cementless, please give details of the surface finish (e.g. grit blasted / sintered bead, with or without HA), including full or partial coating
Modularity
Stem material(s) and manufacturing method(s)
Offset ranges (eg standard, high offset)
Femoral stem length range (minimum to maximum)
Compatible head(s) - materials and manufacturing methods
Femoral head(s) sizes range (minimum to maximum)
Other design features that may affect implant survival
If you answered 'Yes', please explain
Compatible Acetabular cups and liners.
Stem design history
Date of first clinical use
Date of the original CE mark
Date of the latest CE mark renewal (new MDR / UKCA)
Date of first clinical use in UK
Have any design changes been made?
If yes, please give details and dates of modifications | 1
2
3
Please confirm that all data submitted relates to the latest design
Yes
No (please give justification)
Please confirm that the revision rates for all stem versions listed in the product codes falls within the benchmark for which you have applied
(NB You may be asked to provide supporting evidence)
Benchmark claimed (select from dropdown list) | →
Does this product have a current ODEP rating?
Current rating (if applicable) | →
Year rated | →
If known, what is the listing URL (the exact address of the webpage) on the ODEP website? |  | http://www.odep.org.uk/
Beyond Compliance Product
Please confirm a list of UK implanting centres is attached
Is the clinical data submitted for this stem representative of the results of all studies conducted in relation to it?
If 'no', please give details of, and reasons for, data that has been omitted that does not indicate similar or better results than that submitted on the attached Clinical Data Templates
Have any Field Safety Notices (FSNs) or similar notices been issued / served for this product since the current rating was given? If "Yes", please provide further details in the comments box NB! ODEP will make a decision as to whether this information to have a bearing on the submission, but knowlingly failing to declare any FSNs could lead to the rating being removed.
Have RSA studies been conducted on this prosthesis? If "Yes", please attach details of date of study(s), number of patients evaluated, where the study was undertaken and results including time interval between first and second X-Ray NB! ODEP will welcome details of RSA studies as they feel they are important but it is realised that they are not always available and failure to produce RSA studies will not lead to the submission being rejected
Yes
No
Cemented
Cementless
Monobloc head
Modular head
Modular neck
Monobloc head
Modular head
Modular neck
Yes
No
Yes
No
Yes
No
Yes
No
Yes
No
Yes
No
Confirm
Cannot confirm
Yes
No
Yes
No
Yes
No
Yes
No
Yes
No
Yes
No
Yes
No
Yes
No
Yes
No
Yes
No
Yes
No
Yes
No
Yes
No
Yes
No
Yes
No
Yes
No
Yes
No
Yes
No
Yes
No
Yes
No
Yes
No
Yes
No
Yes
No
Yes
No
Collared
Collarless

## Prosthesis Fields - Cup
Hip Prosthesis Data - Cup
Please read the guidance notes before filling out this sheet
Prosthesis details
Manufacturer
UK Distributor (if different)
Cup brand (please name all variants or versions within the Brand)
Product Codes and Sizes |  | Please provide on a separate sheet detailing individual product numbers and descriptions in Excel format
Images of products attached, including all variants Tickbox - yes / No |  | Photo taken against white background, best size 300x300px, resolution 200dpi. This could be used on the ODEP website.
The date (year) of the first ODEP award for this product
Date of this submission
Please confirm for the product codes listed
Data presented relates to use in primary hip replacement only
Material composition is identical
Bone facing surface finish and coating is identical (cementless)
Bone facing design is identical (cementless)
Modular liner mechanism options (cementless)
For modular shells confirm that dual mobility liners have been submitted separately.
If you answer 'no' to any of the above please explain and ensure it is clear what differences exist and the product codes they relate to
Technical design features for the product codes listed
Fixation
Modularity
Dual mobility |  | [Free text]
Constraint
If cementless describe:
Primary coating type(s)
Fixation method(s)
Shell or monobloc - material(s) and manufacturing method(s)
Liner - material(s) and manufacturing method(s)
Differences in the design of liners
Shell/liner locking mechanism
If cemented describe:
Cup material(s) and manufacturing method(s)
Cup design differences
For all cup types -
Acetabular cup or shell / liner(s) sizes range (minimum to maximum)
Internal diameter size ranges (minimum to maximum)
Other design features that may affect implant survival |  | [Free text]
If you answered 'Yes', please explain
Compatible Femoral stem and head(s)
Cup design history
Date of first clinical use
Date of the original CE mark
Date of the latest CE mark renewal (new MDR / UKCA)
Date of first clinical use in UK
Have any design changes been made?
If yes, please give details and dates of modification | 1
2
3
Please confirm that all data submitted relates to the latest design
Yes
No (please give justification)
Please confirm that the revision rates for all design options/features falls within the benchmark for which you have applied
(NB You may be asked to provide supporting evidence)
Benchmark claimed (select from dropdown list) | →
Does this product have a current ODEP rating?
Current rating (if applicable) | →
Year rated | →
If known, what is the listing URL (the exact address of the webpage) on the ODEP website? |  | http://www.odep.org.uk/
Beyond Compliance Product
Please confirm a list of UK implanting centres is attached
Is the clinical data submitted for this cup representative of the results of all studies conducted in relation to it?
If 'no', please give details of, and reasons for, data that has been omitted that does not indicate similar or better results than that submitted on the attached Clinical Data Templates
Have any Field Safety Notices (FSNs) or similar notices been issued / served for this product since the current rating was given? If "Yes", please provide further details in the comments box NB! ODEP will make a decision as to whether this information to have a bearing on the submission, but knowingly failing to declare any FSNs could lead to the rating being removed.
Have RSA studies been conducted on this prosthesis? If "Yes", please attach details of date of study(s), number of patients evaluated, where the study was undertaken and results including time interval between first and second X-Ray NB! ODEP will welcome details of RSA studies as they feel they are important but it is realised that they are not always available and failure to produce RSA studies will not lead to the submission being rejected
Cemented
Cementless
Monobloc
Modular
Yes
No
Confirm
Cannot confirm
Yes
No
Yes
No
Yes
No
Yes
No
Yes
No
Yes
No
Yes
No
Yes
No
Yes
No
Yes
No
Yes
No
Yes
No
Yes
No
Yes
No
Yes
No
Yes
No
Yes
No
Yes
No
Yes
No
Yes
No
Yes
No
Yes
No
Yes
No
Yes
No
Yes
No
Yes
No
Yes
No
Yes
No
Yes
No
Yes
No
Yes
No
Yes
No
Yes
No
Yes
No
Yes
No
Yes
No
Yes
No
Yes
No
Yes
No
Yes
No
Yes
No
Yes
No
Yes
No
Yes
No

## Product image
Please insert an image
Can this image be used on the ODEP website?
If more than one version exists within a brand data being submitted, please include one image of each version
Yes
No

## Product codes
Product code listing
® Please only include the product codes relating to the construct being submitted, to ensure that ratings can be tracked by product code as well as brand.
® This information needs to be provided in the table form below and a separate code for each size (no ranges)
Supplier | Product Brand/Name | Product Description | Product Code | Component | Fixation

## Clinical data sheet 1
Reference Type |  | Reference: |  |  |  | Years | 10 | 3 | 5 | 7 | 10 | 13 | 15
|  | Please insert full details of the reference source. See guidance notes. |  |  |  | Offset | 4
For "Registry data" - the applicant agrees that ODEP can confirm the accurate use of the data with the registry involved | [6 empty] | Criteria |  |  | Meets criteria
For "In-house data" - the applicant confirms that the in-house data used for this submissions complies with ISO 14155:2020 standards | [5 empty] | Criteria - Total Hip Replacement | A* | A | B | 10A* | 10A | 10B |  | Unsatisfactory | Incomplete | Reason
Cohort Prosthesis Details | [14 empty] | FALSE | Product type
Product Type (Stem or Cup) | Cup | [13 empty] | FALSE | Kaplan Meier survival time
STEM brand(s) and variant, if applicable
HEAD brand(s) and variant, if applicable
CUP brand(s) and variant, if applicable | IP CUP (X-LINKED) |  |  |  |  | Minimum number of centres outside development centre(s) | 3 | 3 | 1 | TRUE | TRUE | TRUE | FALSE | FALSE | FALSE | Number of centres
LINER brand(s) and variant, if applicable | [5 empty] | Minimum number of surgeons outside of the development centre | 3 | 3 | 1 | #N/A | #N/A | #N/A | #N/A | #N/A | TRUE | Number of surgeons
[10 empty] | #N/A | #N/A | #N/A | #N/A | #N/A | TRUE
Clinical Study Details | Total Cohort |  |  | Comments: |  | Minimum total cohort | 500 | 500 | 100 | TRUE | TRUE | TRUE | FALSE | FALSE | FALSE | Total number of hips
Number of centres | 8 |  |  | The register do not have information on surgeons, only hospitals. There is 2117 primary operation in 1937 patients. Median age is 74.9 with (min = 14.2, max =96.3). In the fall 2017 the register started collecting PROMs. Patients are examined at the operation and at a one year control. Some results are publshed in our Årsrapport. Diagnoses are based on operations, including bilaterals (n = 2117). The hiarki for the diagnosis is RA>OA>OTHER. Comulative revision rates are calculated at 10 years. X-LINKED was opearted from 2011(only 6 used in 2011). The last cup/stem revisjon was at 10.1 years. For the follow up in years (min = 0 max = 12.9) . |  | Minimum at risk at benchmark time | 400 | 51 | 40 | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | Number ar risk at benchmark time
Number of surgeons implanting | [5 empty] | Max Component upper 95% CI |  |  |  | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | 95% upper confidence limit
Is the study by product developer (designer surgeon)? | [5 empty] | Max Component lower 95% CI |  |  |  | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | 95% lower confidence limit
if yes, please provide details of the study in the comments box
[6 empty] | Max Stem upper 95% CI | 5% | 7% | 100% | TRUE | TRUE | TRUE | TRUE | FALSE
Clinical Study Design (please select from dropdown list) | [5 empty] | Max Stem lower 95% CI | 5% | 7% | 5% | TRUE | TRUE | TRUE | TRUE | FALSE
[6 empty] | Max Cup upper 95% CI | 5% | 7% | 100% | TRUE | TRUE | TRUE | TRUE | FALSE
[6 empty] | Max Cup lower 95% CI | 5% | 7% | 5% | TRUE | TRUE | TRUE | TRUE | FALSE
Patients and Clinical Results (all implants in study)
Total number of patients | 1937 |  |  |  |  | Rating | 10B |  |  | FALSE | FALSE | TRUE | FALSE | FALSE | TRUE | : Number of surgeons missing
Number of males (%) | 523 | 27%
Number of females (%) | 1414 | 73% |  |  |  | Criteria - A* Ratings | 3A* | 5A* | 7A* | 10A* | 13A* | 15A*
Mean age of patients and range (years) | 74.7 | (14,2-96,3) |  |  |  | Minimum number of centres outside development centre(s) | 3 | 3 | 3 | 3 | 3 | 3
[6 empty] | Minimum number of surgeons outside of the development centre | 3 | 3 | 3 | 3 | 3 | 3
Diagnosis: | [5 empty] | Minimum total cohort | 150 | 250 | 350 | 500 | 500 | 500
OA | 1694 |  |  |  |  | Minimum at risk at benchmark time | 150 | 225 | 300 | 400 | 400 | 400
RA | 83 |  |  |  |  | Maximum revision rate ‡ | 3% | 3.5% | 4% | 5% | 6.5% | 8%
Other | 340
[6 empty] | Criteria - A Ratings | 3A | 5A | 7A | 10A | 13A | 15A
Total number of hips in study | 2117 |  |  |  |  | Minimum number of centres and surgeons | 3 | 3 | 3 | 3 | 3 | 3
Number of implants lost due to patients death: | 341 | 16.11% |  |  |  | Minimum total cohort | 150 | 250 | 350 | 500 | 500 | 500
(explain death rate if above 35% in comments box) | [5 empty] | Minimum at risk at benchmark time | 72 | 66 | 60 | 51 | 42 | 40
Number of implants lost to follow-up | 2 | 0.09447% |  |  |  | Maximum revision rate ‡ | 5% | 5.5% | 6% | 7% | 8.5% | 10%
Number of implants revised for any reason | 112 | 5.291% |  |  |  | ‡ The upper 95% confidence interval for KM revision rate (1 - Survival) must be lower than the specified level
Number of implants surviving at end of study | 1660 | 78.41%
Number of patients examined in person | [5 empty] | Criteria - B Ratings | 3B | 5B | 7B | 10B | 13B | 15B
Number of patients examined by questionnaire/telephone review | [5 empty] | Minimum number of centres and surgeons | 1 | 1 | 1 | 1 | 1 | 1
[6 empty] | Minimum total cohort | 100 | 100 | 100 | 100 | 100 | 100
Mean follow-up and range (years) of implants in full cohort | 4.610 | (0-12.9) |  |  |  | Minimum at risk at benchmark time | 40 | 40 | 40 | 40 | 40 | 40
[6 empty] | Maximum value of 95% lower confidence limit for revision rate | 3% | 3.5% | 4% | 5% | 6.5% | 8%
Reasons for revision in full cohort | Number | Percentage
Infection | 38 | 1.795%
Pain | 8 | 0.3779% |  |  |  | Randomised Controlled Trial
Recurrent dislocation | 15 | 0.7085% |  |  |  | Comparative study
Fracture | 5 | 0.2362% |  |  |  | Unselected consecutive cohort
Malposition/malalignment |  | 0% |  |  |  | Selected cohort
Aseptic loosening |  | 0% |  |  |  | National registry
Wear |  | 0% |  |  |  | Hospital Registry
Other |  | 0% |  |  |  | Company database
Other Registry or database (please describe in comments box)
Kaplan-Meier Survivorship
Kaplan Meier survival time in years | 10
Number of hips remaining at risk at Kaplan Meier survival time | 41
| Revision Rate | 95% Confidence Interval
Cumulative Revision Rate (1 - Survival) |  | Lower CI | Upper CI
All failure modes - any component revised (%) | 7.2% | 5.5% | 8.9%
All failure modes - stem revised(%) | 0.4% | 0.1% | 0.8%
All failure modes - cup revised (%) | 1.9% | 1% | 2.9%
Provisionally meets criteria for rating: | 10B: Number of surgeons missing
ODEP-Hip-SubmissionForm_v8.1
Yes
No

## Rating System
Criteria - Total Hip Replacement
Criteria - A* Ratings | 3A* | 5A* | 7A* | 10A* | 13A* | 15A*
Minimum number of centres outside development centre(s) | 3 | 3 | 3 | 3 | 3 | 3
Minium number of surgeons outside of development centre(s) | 3 | 3 | 3 | 3 | 3 | 3
Minimum total cohort | 150 | 250 | 350 | 500 | 500 | 500
Minimum at risk at benchmark time | 150 | 225 | 300 | 400 | 400 | 400
Maximum revision rate ‡ | 3% | 3.5% | 4% | 5% | 6.5% | 8%
Criteria - A Ratings | 3A | 5A | 7A | 10A | 13A | 15A
Minimum number of centres and surgeons | 3 | 3 | 3 | 3 | 3 | 3
Minimum total cohort | 150 | 250 | 350 | 500 | 500 | 500
Minimum at risk at benchmark time | 72 | 66 | 60 | 51 | 42 | 40
Maximum revision rate ‡ | 5% | 5.5% | 6% | 7% | 8.5% | 10%
‡ The upper 95% confidence interval for KM revision rate (1 - Survival) must be lower than the specified level
Criteria - B Ratings | 3B | 5B | 7B | 10B | 13B | 15B
Minimum number of centres and surgeons | 1 | 1 | 1 | 1 | 1 | 1
Minimum total cohort | 100 | 100 | 100 | 100 | 100 | 100
Minimum at risk at benchmark time | 40 | 40 | 40 | 40 | 40 | 40
Maximum value of 95% lower confidence limit for revision rate | 3% | 3.5% | 4% | 5% | 6.5% | 8%
Pre-Entry A* | Pre-Entry
Product launched under Beyond Compliance | Product details supplied to ODEP

## Post Market clinical follow-up
The following information is required for all studies (published or unpublished) initiated by the manufacturer for all products applied for all benchmarks
Prosthesis | Implanting centres | Number of surgeons | Cohort size | Methodology of study

## Implanting Centres
Please use this sheet to list all centres that are implanting the brand family and variant you are submitting
Centre Name (required) | Number implanted over previous 12 month period (Optional)
| Full brand family | Brand Variant under submission